Computed Radiography System

| S. N. | Purchaser’s Specifications |  |
| --- | --- | --- |
| 1 | Computed Radiography System |  |
| 2 | Manufacturer |  |
| 3 | Brand |  |
| 4 | Type Model |  |
| 5 | Country of Origin |  |
| 6 | Manufacturer Authorization Letter |  |
| 7 | Description of Functions Radiography system to replace conventional film/screen based X-ray processing techniques with Photo stimulable Phosphor plate technology to obtain digital X-ray image. |  |
| 8 | Operational Requirements Shall operate on single phase 220-240 v AC supply. Operating environment:12-40c° Humidity:20-80 % RH, non Condensing |  |
| 9 | System Configurations CR system shall consist of: - image reader (CR reader/digitizer):1unit - Processing Server and Workstation:1 unit - Dry Imager:1 unit - Image recording system (cassettes and reading plaes) |  |
| 10 | Technical Specification The system shall be compatible for general Radiography application. Image reader ( CR reader / digitizer): - should have scanning speed of >50 sheet per hours for the largest film size of 14"*17" - The initial film preview time shall be less than 50 seconds per cassette size of 14"*17" - The Gray scale resolution for CR reader/ digitizer must have a minimum resolution of 12 bits/pixel for image sent to CR processing station. - capability for quick check of the image and exam data of at least the last 4 imaging plates scanned at the x-ray room. - The Cr reader/digitizer must be able to handle phosphor image plate. |  |
| 11 | Processing Server and Workstation - The processing station must have at least 4 GB RAM, 500GB HDD and ≥19 inch branded high resolution LED Monitor. - Must be possible to take print outs directly .Mechanism for printing multiple image in one film, with possibility of slide and true size printing. - Must have a protocol for verifying the connectivity status of configured image destinations. - It should provide display of acquired image with details of demographics viz. patient / study listing for ease access. - The server must provide full amount of post processing features viz. geometric corrections, window level algorithms, annotations like markers, predefined text ,drawing lines and geometrical shapes , multi scale image processing , measuring distance and angle, shuttering ,histograms, zoom, grey scale reversal, edge enhancement, noise reduction ,indication of grey scale saturation level ,latitude reduction, angular rotation etc. - It should be able to send DICOM image to DICOM workstation or PACS without loss of information. - Exporting image in BMP, JPEG and AVI formats on DICOM and NON-DICOM. - Should be able to store image on external device viz. CD/DVD or pen drive etc. - Must have CD and DVD read/write device. |  |
| 12 | Dry Laser Printer - Must be a Laser image printer with a resolution of not less than 500 dpi or more. - The system must be able to print >50 films per hours of the largest size. - The system must have contrast resolution of 14 bits/ pixel or more .The system must have at least two online film sizes tray and should be capable of printing any of the 8"*10",10"*,12",14"*17" films. - The imager should support daylight loading of films. - It must be DICOM compatible allowing multiple modalities to be connected at a time. |  |
| 13 | Accessories, Spare Part and Consumables Accessories: - IP/Cassettes of all Sizes: - 8"*10": 01 pc - 14"*17":01pc - 10"*12":01pc Should provide the list of important spare parts and accessories with their part numbers and costing. All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication material, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 14 | Operating Environment The system offered shall be designed to be stored and to operate normally under condition of purchaser's country. The conditions included power supply, climate, temperature, humidity, climate temperature, humidity etc. - Power Supply 220-240 VAC,50 HZ Single phase fitted with appropriate plug .The power cable must be at least 3 meter on length. |  |
| 15 | Standard and Safety Requirements Must submit ISO13485 : 2003/ AC:2007 for Medical Devices AND European CE or USFDA Approved product certified. Electrical safety conforms to standards for electrical safety IEC 60601-1 General requirement for electrical safety of Medical Equipment. |  |
| 16 | User Training Must provide user’s training including how to use and maintain the equipment. |  |
| 17 | Warranty Comprehensive warranty for 2 years after acceptance. |  |
| 18 | Maintenance Service During Warranty Period During warranty period supplier must be ensure preventive maintenance and corrective / breakdown maintenance whenever required. |  |
| 19 | Installation Commissioning The bidder must be arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail |  |
| 20 | Documentation - User (Operating) manual in English. - Original Brochure must be submitted |  |