ECG Machine, Portable (12 Channel)

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | ECG Machine, Portable (12 Channel) |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | ECG Machine is primary equipment to record ECG Signal in various configurations. |  |
| 2 | Operational Requirements |  |
| 2.1 | Portable digital ECG machine must be able to acquire all 12 Leads simultaneously. |  |
| 3 | System Configuration |  |
| 3.1 | Portable digital ECG machine with complete accessories |  |
| 4 | Technical Specifications |  |
| 4.1 | Simultaneous recording of 12 standard leads: aVR, aVL, aVF, I, II, III and V1-6 pre-cordials. |  |
| 4.2 | Should have built in large memory for data storage. |  |
| 4.3 | Should have Splash-resistant alphanumeric function keys/buttons. |  |
| 4.4 | Should have Filter setting for line-frequency (50 or 60Hz) and tremor. |  |
| 4.5 | Continuous check on the quality of electrodes connection, audio visual alert on loss of signal |  |
| 4.6 | Should have protected defibrillation protection function for operation. |  |
| 4.7 | LCD display, approximately: 7” with high resolution touch screen. Display shows ECG-curves, heart rate, patient name and ID, time, speed and filter setting. |  |
| 4.8 | Front panel provides indication of system and battery status, electrode connection and paper. |  |
| 4.9 | Built-in high-resolution thermal printer, with automatic and manual print-out mode. |  |
| 4.10 | Number of channels printed is user selectable: 3, 6 or 12. |  |
| 4.11 | Paper speed, user adjustable: 5, 6.25, 12.5 25 and 50mm/sec. |  |
| 4.12 | Sensitivity, automatic or user selectable: 5, 10 and 20mm/mV. |  |
| 4.13 | Data interface: RS232 or equivalent Self-test is performed each time the device is switched on. |  |
| 4.14 | Should have ECG signal input sampling frequency of approx. 32kHz. |  |
| 4.15 | Should have rechargeable lithium battery integrated in device with at least 2hours of backup with printing. |  |
| 4.16 | Should have CMRR (Should have CMRR (common mode rejection) 100dB or more for good diagnostic performance. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Patient cable-1 no. - Reusable chest electrodes, suction ball-type- 6 nos. - Extremity clamp electrodes, reusable- 4 nos. - Recording thermal paper- 1Rolls. - Bottles of electrode gel, approximately 250ml- 2 nos. |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220–240V AC, 50Hz fitted with appropriate plug type. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
| 7.2 | CE (93/42 EEC Directives) and USFDA approved product certificate. |  |
| 7.3 | Shall meet IEC-60601-1-2 General Requirements and IEC-60601-2-25 Safety of Electrocardiograms. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | Supplier must accomplish proper installation and commissioning of the equipment on site. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English |  |
| 12.2 | Service (Technical / Maintenance) manual in English |  |