LED View Box (Single Film)

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | LED View Box (Single Film) |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | View box used for viewing the images of X-ray, CT/MRI at healthcare facilities. |
| 2 | Operational Requirements |
| 2.1 | Single film LED view box, operates on mains electric supply. |
| 3 | System Configuration |
| 3.1 | LED View Box (Single Film), complete unit. |
| 4 | Technical Specifications |
| 4.1 | Ultra slim design. |
| 4.2 | LED backlit and shall have separate on/off function with separate rotary continuous adjustable brightness control at the bottom of panel for convenient operation. |
| 4.3 | It shall have fully electronic continuous brightness control with adjustment range approx. up to 90%. |
| 4.4 | Shall have no lag period in intensity modulation. |
| 4.5 | Front sheet shall be made of polycarbonate or acrylic with antiglare. |
| 4.6 | Shall have sturdy film clamping mechanism with automatic sensor induced on/off system. |
| 4.7 | Illumination: High bright white LEDs. |
| 4.8 | It shall have homogeneous illumination and shall have luminance of more than 1200 cd/m2. |
| 4.9 | LED light source shall have at least 20000 hours of operation. |
| 4.10 | Shall be able to hold one full large size CT/MR films at a time with film has maximum size of 43cmX 35cm. |
| 5 | System Configuration Accessories, spares and consumables |
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer. Bidders shall specify, in a separate Excel worksheet, the quantity and details of any items included in this offer which have not been specified in this Technical Specifications Form. |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND |
| 7.2 | CE or USFDA approved product certificate. |
| 8 | User Training |
| 8.1 | Not applicable. |
| 9 | Warranty |
| 9.1 | Warranty for 1 year after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | Standard warranty conditions are applicable. |
| 11 | Installation and Commissioning |
| 11.1 | Must supply preassembled unit, ready to use. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part number and costing. |