# X-Ray Machine 500 mA, Fixed Height Table

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | X-Ray Machine 500 mA, Fixed Height Table |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Functions |
| 1.1 | A general purpose X-ray machine for routine X-ray examinations at healthcare facilities. |
| 2 | Operational Requirements |
| 2.1 | It shall be suitable to be used for adult and paediatric patients in general Radiography examination. |
| 3 | System Configurations |
| 3.1 | X-ray Generator,1 unit |
| 3.2 | X-Ray tube & tube support system, 1 unit |
| 3.3 | Radiographic patient table, fixed height1 unit |
| 3.4 | Floor mounted bucky stand, 1 unit |
| 4 | Technical Specifications |
| I | X-ray Generator: Bidder shall indicate brand and model information here and provide technical data document for X-ray generator offered |
| 4.1 | Microprocessor based, high frequency inverter generator, the generator shall have at least 50kHz. |
| 4.2 | Generator Output: not less than 35 kW (500mA at 100kV) |
| 4.3 | Radiographic voltage: 40 kV to 125 KV, in 1kV step. |
| 4.4 | Radiographic current: 10 to 500mA |
| 4.5 | Exposure time: 0.001sec (1msec) - 10sec |
| 4.6 | Anatomical Programmable Radiographic mode shall be available. |
| 4.7 | Manual & automatic exposure control and automatic brightness control shall be available. |
| 4.8 | Shall come with overload protection device. |
| 4.9 | Power supply: 3 phase, 380-415V 50/60Hz |
| II | X-Ray Tube: |
| 4.10 | X-ray tube rotating: +/-120°. |
| 4.11 | Large focus not more than 1.2 mm. |
| 4.12 | Small focus not more than 0.6 mm. |
| 4.13 | Maximum tube voltage 150 KV. Maximum tube output shall match with the generator output of not less than 35 KW. |
| 4.14 | Filtration: min 2.5mm Al equivalent. |
| 4.15 | Cooling method passive or forced air and/or oil cooling. |
| 4.16 | Anode rotating speed: More than 3000 rpm. |
| 4.17 | Anode heat capacity shall not be less than 300 KHU. |
| III | Radiography Patient Table: |
| 4.18 | Radiography table shall be fixed height, 4-way floating top type with foot switch control. |
| 4.19 | Come with grid and cassette tray, with grid ratio: not less than 10:1. Grid line number: 40 line/cm. Focus distance: 115cm. |
| 4.20 | Cassette size: accept all sizes from cassette 5”x7” cm to 14”x17” type. |
| 4.21 | Radiography table shall be fixed height of about 60cm. |
| 4.22 | Table top to film distance: approx. 6cm. |
| 4.23 | Table top transverse movement : approx. ±14cm. |
| 4.24 | Table longitudinal movement : approx. ± 29cm. |
| 4.25 | Table top dimension: approx. 2000 mm x 800 mm. |
| 4.26 | Shall come with a three-field AEC. |
| 4.27 | Table movement arrested by electromagnetic brakes. |
| IV | Floor Mounted Bucky Stand: |
| 4.28 | Vertical travel: from 460-1700mm or in the range. |
| 4.29 | Moving Grid with Grid ratio not less than 10:1. Grid line number: 40 lines/cm. |
| 4.30 | Shall come with Automatic Exposure Control for vertical bucky exposures. |
| 4.31 | Cassette size: accept all sizes from 5”x7” to 14”x17”. |
| 4.32 | Movement arrested by electromagnetic brakes. |
| V | Floor Mounted Tube Stand: |
| 4.33 | Longitudinal travel: approx. 1750mm. |
| 4.34 | Vertical travel: from 630 -1850mm or in the range. |
| 4.35 | Movement arrested by electromagnetic brakes. |
| 4.36 | Rotation of tube arm around vertical axis: 1800; lockable at 00 to +/- 900. |
| VI | Collimator: |
| 4.37 | Manually adjustable. |
| 4.38 | Manually selectable filters. |
| 4.39 | Light localizer with timer controlled light. |
| 4.40 | Built-in light switch should be provided. |
| 4.41 | Turning angle should be min +/- 45 degree. |
| 4.42 | Halogen lamp. |
| VII | Control Console: |
| 4.43 | Digital Display. |
| 4.44 | Minimum 3 Point Exposure Technique. |
| 4.45 | Status display, error display. |
| 4.46 | Shall have area dose product determination and display |
| 4.47 | Shall come with radiography remote control in control room |
| 5 | Accessories, Spare Parts and Consumables |
| 5.1 | Accessories: - Lead apron, light weight with Lead equivalence 2mm-02 nos. |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 380-415VAC 3 phase 50Hz fitted with appropriate plug for X-ray generator fitted with appropriate plug for other units. The power cable must be at least 3 metres in length. |
| 7 | Standards & Safety Requirements |
| 7.1 | Must submit ISO 13485:2003/AC: 2007 AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Shall meet: - IEC 60601-1-3 - Part 1: General Requirements for safety - Collateral Standard: General Requirements for Radiation Protection in Diagnostic X-Ray Equipment. - IEC 60601-2-7 - Part 2-7: Particular Requirements for the Safety of High-Voltage Generators of Diagnostic X-Ray Generators. |
| 8 | User Training |
| 8.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users. |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years from acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |