Autoclave, Horizontal 75 litres

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Autoclave, Horizontal |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
|  | # Autoclaves are required for sterilizing objects in high temperature and high pressure steam. |  |
| 2 | Operational Requirements |  |
|  | Horizontal electrically heated autoclave is required. |  |
|  | Shall be used with distilled water. |  |
|  | Shall be mains power operated. |  |
| 3 | System Configuration |  |
| 3.1 | Autoclave, Horizontal 75L with complete accessories. |  |
| 4 | Technical Specifications |  |
|  | Dimensions: - Chamber Diameter: 350 mm approx. or better - Chamber Depth: 600mm approx. or better |  |
|  | Industrial immersion type water heater to generate steam within a reasonable period of time on 3 phase 50 HZ ac supply. |  |
|  | Operating pressure: 1.2kg/cm2 |  |
|  | Operating temperature: 121-134 °C. |  |
|  | Material of Construction: - Triple walled construction. - Piping, fittings, Inner chamber, Jacket, Door: SS 304/316 good quality. - Outer Chamber: SS 304 (insulated properly). - Steam generator: Non corrosive SS /Chromium plated Brass. - Heater Plate: Brass/Stainless steel. |  |
|  | Shall have temperature, pressure and vacuum gauges. |  |
|  | Load bearing capacity 3KW or more. |  |
|  | Heating: Direct steam, provided from an Electrically-heated built-in generator. |  |
|  | Digital display – temp. and time |  |
|  | Stainless steel flush mounting. |  |
|  | Blue water/ Vacuum leak detection tests for vials etc. |  |
|  | Carriages, trolleys, trays, baskets etc. for efficient transportation of goods pre and post sterilization. |  |
|  | Centralised Control Point: - Multi – Port valve control for the entire cycle of operation. - Permit two exhaust speed ; Fast and Slow - Easy to read compound gauge for the chamber and pressure gauge for the jacket |  |
|  | Safety Features: - Self-locking safety door – cannot be opened when chamber is under pressure. - Gauge glass with safety valves. - Electrically operated pressure-controlled device. - Heat resistant door silicon gasket for steam tight sealing. - Low water protection system to protect the electric heaters. - Drain temperature gauge – digital temp. indicators / controllers. - Vacuum breakers - Safety valve for the jacket. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Spare heating element- 2 set - Indicator tape – 5 pcs - Autoclave drums- 4 - Bowie dick test pack - 5 |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | # The system offered must be designed to store and be operated normally under the condition of the purchaser’s Country. The conditions include Climate, temperature and relative humidity. |  |
| 6.2 | Power supply: 380-440 V (3 Phase), 50Hz fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |  |
| 7 | Standards and Safety Requirements |  |
|  | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
|  | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
|  | Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |  |
| 10 | Maintenance Service during Warranty Period |  |
| 10.1 | During warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |