Autoclave, Vertical

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Autoclave, Vertical |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
|  | Autoclaves are required for sterilizing objects in high temperature and high pressure steam. |  |
| 2 | Operational Requirements |  |
|  | Electrically heated vertical steam sterilizer. |  |
|  | Shall be used with distilled water. |  |
|  | Shall be mains power operated. |  |
| 3 | System Configuration |  |
| 3.1 | Autoclave, Vertical with complete accessories. |  |
| 4 | Technical Specifications |  |
|  | Triple walled construction with dual door. Chamber, door, door frame, bolts made of corrosion-resistant material and able to prevent stress cracking. |  |
|  | Operating pressure range: 15-30 psi adjustable. |  |
|  | Operating temperature range: 121-134 C |  |
|  | Pressure control switch with digital display. |  |
|  | Outer and inner chamber made of stainless steel. |  |
|  | Inner chamber made of at least 18 SWG SS sheet |  |
|  | Inner chamber size approx.. 400X400x600 mm or cylindrical with a dia. of 400mm and a height of 600 mm. |  |
|  | Stainless steel Steam jacket insulated with high grade glass wool. |  |
|  | Shall have temperature, pressure and vacuum gauges. |  |
|  | Shall have: Steam traps, vacuum driers, water level indicator on steam generator, safety devices, pressure switch and safety valve, self-locking of door when chamber is under pressure, vacuum breaker for jacket. |  |
|  | Water level indicator with automatic low water level cut off device. |  |
|  | Joint less silicon gasket |  |
|  | Water inlet and drain valves. |  |
|  | Safety Features: - Self-locking safety door – cannot be opened when chamber is under pressure. - Gauge glass with safety valves. - Heat resistant door silicon gasket for steam tight sealing. - Pressure switches to economise power consumption. - Low water protection system to protect the electric heaters. - Drain temperature gauge – digital temp. indicators / controllers. - Vacuum breakers - Safety valve for the jacket. - Hydraulically tested to 2 times the designed working pressure for jacket and 1.5 times for chamber. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Spare heating element- 2 set - Three stainless steel wire baskets; so that one/two suitably fit into the autoclave at the same time. - Three stainless steel buckets with lid; so that one/two suitably fit into the autoclave at the same time. |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase or 380-400V AC 50 Hz Three phase, 50Hz. fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |  |
| 7 | Standards and Safety Requirements |  |
|  | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
|  | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
|  | Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |  |
| 10 | Maintenance Service during Warranty Period |  |
| 10.1 | During warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English, |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |