Foetal Heart Ultrasound-Detector

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Doppler Foetal Heart Detector with Rechargeable Battery |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
|  | Doppler foetal heart detector is a hand-held ultrasound transducer used to detect the heartbeat of a foetus for prenatal care. |  |
| 2 | Operational Requirements |  |
|  | Shall be portable/handheld, lightweight and easy to carry. |  |
| 3 | System Configuration |  |
| 3.1 | Doppler, Foetal Heart Detector, complete with accessories. |  |
| 4 | Technical Specifications |  |
| 4.1 | Doppler based foetal heart rate detector with amplifier loudspeaker. |  |
| 4.2 | Transducer frequency, approx.: 2 MHz. and output power shall not be more than 10mW/cm2. Must be capable of detecting foetal heart rate within the range of 50 to 240bpm |  |
| 4.3 | Transducer probe with fixed wire connection to the main unit, length approximately 35cm. |  |
| 4.4 | Detector diameter approximately 20mm. |  |
| 4.5 | Self-test is performed each time the device is switched on. |  |
| 4.6 | Large LCD/TFT display shows foetal heart rate (FHR) in beats per minute (bpm), pulse indicator, sound volume level, battery indicator. |  |
| 4.7 | System shall report operational status, malfunctions and low battery with audio-visual alerts. |  |
| 4.8 | Built-in loudspeaker with volume adjustment. |  |
| 4.9 | Advanced noise suppression system assures quality diagnostic sound. |  |
| 4.10 | Operates on inbuilt rechargeable batteries with autonomy of 10 hours or more continuous operation. Bidder to specify the number of batteries to be supplied, and details of hours of examinations on fully charged battery. |  |
| 4.10 | Battery charger with AC adaptor shall be provided. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - 2 x Tubes of ultrasound gel, approximately 350ml - 1 x Set of spare batteries - 1 x Soft carry bag easy to clean |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power input: 220 – 240 VAC, 50Hz for battery charging. |  |
| 7 | Standards and Safety Requirements |  |
|  | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
|  | CE (93/42 EEC Directives) or USFDA approved product certificate will be prioritized. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |  |
| 10 | Maintenance Service during Warranty Period |  |
| 10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | Supplier must accomplish proper commissioning of equipment onsite. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |