Ice Lined Refrigerator (ILR) 71 ltr

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Ice Lined Refrigerator (ILR) |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | Ice-lined refrigerators are used to store vaccine and maintain temperatures +2°C to +8°C and able to run on less than 5 hours electricity per day. |  |
| 2 | Operational Requirements |  |
|  | Only one vaccine storage compartment. |  |
| 3 | System Configuration |  |
| 3.1 | Ice Lined Refrigerator (ILR), single compartment with suitable voltage stabiliser (As specified in 6.3). |  |
| 4 | Technical Specifications |  |
| 4.1 | Vaccine storage capacity: Minimum 45 litres with baskets and shelves in place. |  |
| 4.2 | Refrigerator compartment internal Gross Volume: Maximum 71 litres. |  |
| 4.3 | Construction: - Corrosion resistance: Internal and external cabinet, lid and frame protected against corrosion. Bidder to specify the materials used for construction of internal & external cabinet. - CFC free insulation. - Door with lock and handle. |  |
| 4.4 | Refrigerant: The refrigerator shall utilize CFC (Chlorofluorocarbon) free refrigerants |  |
| 4.5 | Holdover time: The refrigerator shall maintain +2 oC to +8oC for minimum 24 hours during power cuts at the ambient temperature +43oC. Preference will be given to longer holdover time. |  |
| 4.6 | Compressor: - Sealed heavy duty type. - Starts and operates on electricity supply from 165 to 255 volts. |  |
| 4.7 | Control Panel: The thermostat, thermometer and other visual displays like On/Off switch, power indicator must be positioned. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Vaccine Storage baskets (wires type) or shelves- minimum 2per unit. - External Reading Thermometer (mercury free) - one piece per unit range of -50 to +50 0C. - Integrated Digital Temperature display LCD/LED-01 (for displaying temperature inside the system and shall operate either electrically or solar power). |  |
| 5.2 | Spare parts: Shall be included complete sets of spare parts in every 10 units (like compressor, indicator, lights, thermostats etc.) |  |
| 6 | Operating Environment |  |
| 6.1 | The product offered shall be designed to be stored and to operate normally under Nepal’s Power Supply, Climate, Temperature (+5 °C to +45 °C), low and high Humidity conditions. |  |
| 6.2 | Unit Power supply: 220-240VAC, 50Hz fitted with appropriate plug. The power cable must be minimum 3 metres long. Power consumption: Preference will be given to less power consumption. |  |
| 6.3 | Voltage Stabiliser: Voltage stabiliser unit to fit the refrigerator shall be included. Stabilizer to have the capacity of minimum 0.5 KVA, normally operating input voltage of 110-270V. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Shall meet WHO/UNICEF Standard E3/RF.3.for Ice Lined Refrigerators or shall be WHO PQS certified product. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Documentation |  |
| 11.1 | User (Operating) manual in English. |  |
| 11.2 | Service (Technical / Maintenance) manual in English. |  |