Light, Operation Theatre, Mobile

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Light, Operation Theatre, Mobile |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type / Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | Mobile operating light is required for carrying out operations in an emergency environment and the system can be moved from place to place in hospitals. |  |
| 2 | Operational Requirements |  |
|  | Shall operate on mains electric supply as well as on battery. |  |
| 3 | System Configuration |  |
| 3.1 | Operating Light Mobile with single light head, moveable on casters and with all standard accessories. |  |
| 4 | Technical Specifications |  |
| 4.1 | The light shall be designed with good counterbalance mechanism in order to ensure stability of light head in all positions |  |
| 4.2 | Shall have single light head. |  |
| 4.3 | Light head shall not be greater than 400mm diameter |  |
| 4.4 | Number of bulb: 1 main bulb halogen 24V 150W with 1 backup bulb with auto-switching or multi-bulbs light head not more than 3 bulbs of halogen 12V 50W. |  |
| 4.5 | Light intensity range, shall not be less than 80,000 lux at 1 meter distance from light source. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.6 | Light temperature, between 4000 - 4500 K. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.7 | Shall have colour rendering index in between 92 - 94. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.8 | Shall have 99% heat filtrating. |  |
| 4.9 | Temperature increase at head shall not be more than 2 oC. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.10 | Temperature increase on operating field shall not be more than 15 oC. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.11 | Radiant heat energy (beam temperature) shall not exceed 25,000 microwatts per square centimetre measured 1m from the light source. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.12 | The light offered shall have safety designed to prevent patient from burns, especially during the ophthalmic procedures. The light offered shall be certified safe to be used under ophthalmic procedures. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF. |  |
| 4.13 | Working distance range (focal length): 70 - 130cm. |  |
| 4.14 | Depth of field with focused light: > 60cm. |  |
| 4.15 | Luminance field size: 15 -25cm diameter, adjustable. |  |
| 4.16 | Shall have a control to regulate light intensity and to switch on the unit. |  |
| 4.17 | Shall have an On/Off switch at lamp head. |  |
| 4.18 | Shall come with continuous dimmer, continuous focus adjustment, continuous field adjustment. |  |
| 4.19 | Sterilizeable handle to regulate light field size, 2 pcs. |  |
| 4.20 | Vertical adjustment shall not be less than 115cm. |  |
| 4.21 | Rotation 360 degree. |  |
| 4.22 | Come with Ophthalmic procedures safe light bulbs with a minimum of 1000 hours lifespan. |  |
| 4.23 | Interference suppressed complies with VDE 0875 or equivalent. |  |
| 4.24 | Transformer and operating elements shall be integrated in light head housing. |  |
| 4.25 | Mobile Stand: Shall be based on light weight easily moveable stable support with at least 4 casters with locking counter balance mechanism in order to ensure stability of light head in all positions and with swivel arm. Caster must be medical chemical resistant. |  |
| 4.26 | Battery: - Lithium ion built in rechargeable batteries with capacity sufficient for operating in battery mode (fully charged) for minimum of 3 hours. - Shall include a built-in automatic battery charger with proper protection against battery damage. - Shall include battery power (charge) indicator. Bidder to specify the number of batteries to be supplied. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - 1 x spare halogen bulb. - 1 x spare set of fuses. |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
| 7.3 | Electrical safety conforms to standards for electrical safety IEC 60601-1 General requirement for Electrical safety of Medical Equipment. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |