Haematology Analyser

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Haematology Analyser, 22 parameters |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
|  | Description of Function |
|  | Automated haematology analyser used for determination of a range of parameters in blood analysis. |
|  | Operational Requirements |
|  | Automatic blood cell counter that measures minimum 22 parameters including 3part differential of WBC is required |
|  | System Configuration |
|  | Haematology Analyser, 22 parameter, complete unit with all standard reagents, consumables and accessories. |
|  | Technical Specifications |
|  | Measurement principal: electrical impedance method for WBC, WBC Diff, RBC, MCV, PLT and Cyanide Free Colorimetric method for HGB |
|  | Sample size: Aprox.10ul. |
|  | Throughput: 60 samples per hour or more. |
|  | Determination: WBC, RBC, HGB HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT. MPV , PCT PDW –CV, PDW-SD, P-LCR, P-LCC,LYM#, LYM%, MID #, MID%, Gran# and Gran % 3 Histograms of WBC, RBC & PLT |
|  | Dual counting chamber: WBC Chamber~ 100um and RBC/PLT~70um |
|  | Accuracy & Precision (CV)：WBC ≤ 2.0%, RBC ≤ 1.5%,HGB ≤ 1.5%,MCV ≤ 0.5%, PLT ≤ 4% |
|  | Calibration: Manual and Automatic modes. |
|  | Linearity Range for Measured Parameter WBC : 0 to 300 x 103 /L RBC : 0 to 8.0 x 106/L HGB : 0 to 25g/dl PLT : 0 to 3000 x 103/L HCT : 0 to 65% |
|  | Shall have self-test capability and Automatic Flushing System to remove blockage. |
|  | Display: LCD type of 8” or more touch screen with facility of assistance messages and erroe |
|  | Shall have facility of sample ID, date and time are reported with test results |
|  | Shall come with dedicated data analysis and data management software |
|  | Shall have printing facility through inbuilt thermal printer, paper width of 55mm or more and Various printout formats |
|  | Shall have built-in RS232, USB 2.0 or equivalent port for allowing data transfer. |
|  | Reagent: Machine Should operate with two (Diluent and Lyse) or three (Diluent, cleaner and Lyse) reagents for testing procedure. Two reagent system will be preferred for cost effectiveness and Reagent Cost per test should be declared separately |
|  | Accessories, spares and consumables |
|  | All standards accessories, consumables and parts required to operate the equipment, including all standard tools and clearing and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above) |
|  | Operating Environment |
|  | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
|  | Power supply: 220-240VAC, 50Hz fitted with appropriate plug type D (3 pins). The power cable must be minimum 3 meters long. |
|  | Standards and Safety Requirements |
|  | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
|  | CE (93/42 EEC Directives) or USFDA approved product certificate. |
|  | Shall meet IEC 61010-2-081 safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-081: Particular requirements for automatic and semi-automatic laboratory equipment for analysis and other purposes. |
|  | User Training |
|  | Must provide user training (including how to use and maintain the equipment). |
|  | Warranty |
|  | Comprehensive warranty for 1 years from date of installation. |
|  | After warranty period, supplier should supply spare parts for minimum 5 years. (to be declared with a letter of commitment in the name of purchaser) |
|  | Maintenance Service During Warranty Period |
|  | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required. |
|  | Installation and Commissioning |
|  | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
|  | Documentation |
|  | User (Operating) manual in English. |
|  | Service (Technical / Maintenance) manual in English. |
|  | Certificate of calibration and inspection from factory. |