Semi-Automated Bio-Chemistry Analyser

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Semi-Automated Bio-Chemistry Analyser |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | The Semi-automated Bio-chemistry Analyser measures biochemical indexes by analyzing blood and other body fluid, then combines with other clinical information, to help diagnose disease, evaluate organs function, identify disease gene and determine the norm for future therapy. |
| 2 | Operational Requirements |
| 2.1 | Semi automated Chemistry Analyser with built in thermal printer. |
| 2.2 | User programmable memory for up to 150 test and minimum 10,000 sample test results. |
| 3 | System Configuration |
| 3.1 | Semi automated chemistry Analyser with inbuilt thermal printer data processor & LCD display, touch screen of at least 7” |
| 4 | Technical Specifications |
| 4.1 | Light Source : Quartz Halogen Lamp |
| 4.2 | Photo detector: Silicon based or equivalent (range approx. 300-850nm) |
| 4.3 | Wavelength Range: Approx. 340 – 600 nm. |
| 4.4 | Wavelength Selection: Automatic |
| 4.5 | Photometric Range: Approx. 0 to 4.0 Absorbance. |
| 4.6 | Calculation Modes: Should have at least 5 calculation modes like - Absorbance/concentration - End point - Two points - Kinetic - biochromatic |
| 4.7 | - Incubation Time 1 to 999 second. - Interval Time 3 to 999 second |
| 4.8 | Quality Control - At least 2 controls per test. |
| 4.9 | Calibration method- should have facility of calibration of at least factor, calibrator, multipoint calibration. |
| 4.10 | Flow Cell- Temperature control by Peltier element with ambient temperature if 25, 30,37 degree Celsius. |
| 4.11 | Incubator- Incubator block with 37 degree Celsius with PID controlled. |
| 5 | Accessories, spares and consumables |
| 5.1 | Accessories: - Trial kits for various parameters, multi-calibrators and multi controls. -01 set |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Shall meet IEC 61010-2-081safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-081: Particular requirements for automatic and semi-automatic laboratory equipment for analysis and other purposes. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years from acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure planned preventive maintenance (PPM) and corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English |
| 12.2 | Service (Technical / Maintenance) manual in English |
| 12.3 | List of important spare parts and accessories with their part number and costing. |