Centrifuge – Desktop 8 tubes 10-15 ml

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Centrifuge – Desktop 8 tubes 10-15 ml |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Centrifuges are required in the Laboratory to separate various components of Blood and any other liquid sample for analysis |
| 2 | Operational Requirements |
| 2.1 | Clinical electrical centrifuge to be used at blood transfusion centres for routine centrifuging tests. The units must be fitted with resiliently mounted motor for vibration free performance – preferably built in 5 speed regulator. |
| 3 | System Configuration |
| 3.1 | Centrifuge – Desktop 8 tubes 10-15 ml |
| 4 | Technical Specifications |
| 4.1 | Microprocessor controlled, swing out rotor with buckets. |
| 4.2 | Facilities, adaptors and accessories for 8 tubes of 10-15 ml |
| 4.3 | Speed 1000-5000 RPM; Max. RCF 1600 g |
| 4.4 | Approx. Size: W300 mm x D300 mm x H 275 mm |
| 4.5 | Quiet operation and low vibration |
| 4.6 | Speed and imbalance regulator plus lid lock |
| 4.7 | Power switch ON/OFF, and timer control knob |
| 4.8 | Suitable to work on 220-240 Volts, single phase 50-60 Hz AC supply; minimum of 3 meters mains cable with earth provision (three pin plug fitted) |
| 5 | Accessories, Spares and Consumables |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug type D (3 pins). The power cable must be minimum 3 meters long. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Must comply with IEC 61010-2-020: Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-020: Particular requirements for laboratory centrifuges. |
| 8 | User Training |
| 8.1 | Not applicable |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | Must supply preassembled unit, ready to use. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English |
| 12.2 | Service (Technical / Maintenance) manual in English |
| 12.3 | List of important spare parts and accessories with their part number and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |