Hot Air Oven (Small)

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Hot Air Oven (Small) |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Hot Air Oven is required for heating a sample under controlled conditions. |
| 2 | Operational Requirements |
| 2.1 | Microprocessor based system. |
| 3 | System Configuration |
| 3.1 | Hot Air Oven (Small) with complete accessories |
| 4 | Technical Specifications |
| 4.1 | Must be made of double walled chamber, Stainless Steel SS 304 grade. |
| 4.2 | Must provide with three heating elements on three sides of the equipment for uniform temperature on all shelves. |
| 4.3 | Door gaskets shall be made of Silicon. |
| 4.4 | Shall have variable microprocessor based digital temperature controller with digital display. |
| 4.5 | Must have a minimum chamber size of 300mm (L) x 300mm (B) x 300mm (H) with 2 stainless steel perforated trays. |
| 4.6 | Shall have provision of air ventilations. |
| 4.7 | Temperature variation +/- 1 ℃ |
| 4.8 | Temperature Range: Ambient to 250 ℃. |
| 5 | Accessories, spares and consumables |
| 5.1 | Accessories: - Thermometer-01 no. - Door gasket- 01 no. |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered must be designed to store and be operated normally under the condition of the purchaser’s Country. The conditions include Power supply, Climate, temperature and relative humidity. |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs. The power cable must be at least 3 metres long. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Must be compliant with IEC 61010-1 :( or any international equivalent e.g. EN/UL 61010) covering safety requirements for electrical equipment for measurement control and laboratory use. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |