Neubauer Counting Chamber

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Neubauer Counting Chamber |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
| 1 | Description of Function |
|  | Used to count blood cells. |
| 2 | Operational Requirements |
|  | Chamber: Improved Neubauer counting Chamber Specialized Microscopic slide |
| 3 | System Configuration |
| 3.1 | Neubauer Counting Chamber with complete accessories. |
| 4 | Technical Specifications |
|  | Inside cell of Chamber: An area of 9 Sq. mm (3x3 sq. mm) divided into 9 squares of one mm side by means of triple lines. The central small square is further divided by triple lines into 5x5 = 25 smaller squares each with side 1/5 of an mm. Each of the squares is further divided into 4 x 4 = 16 smallest squares with side equal to one twentieth of an mm. Thus Dimension of individual cell of 0.05 x 0.05 = 0.0025 mm2 |
|  | Gap between cover: 0.1 mm. Slip and grid area |
|  | Pipettes: RBC and WBC pipettes |
|  | Accessories: Rubber suckers for RBC & WBC pipettes |
|  | Cover slips: 22 x 25 mm one box |
|  | Wooden Box: Inside dimensions approx. 100 x 175 x 30 mm |
| 5 | Accessories, spares and consumables |
| 5.1 | Accessories: - Additional two RBC & WBC pipettes each with every Neubauer Chamber. |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |
| 6.3 | Voltage corrector/stabilizer of appropriate ratings meeting international standards.( Input 160-260 V and output 220-240 V and 50 Hz) |
| 7 | Standards and Safety Requirements |
|  | Must submit ISO 9001 or ISO13485:2003/AC:2007 AND |
|  | CE (EEC Directives) or USFDA approved product certificate. |
|  | Shall meet IEC 61010-1 for electrical safety requirements. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |
| 10 | Maintenance Service during Warranty Period |
| 10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English |
| 12.2 | Service (Technical / Maintenance) manual in English |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |