Refrigerator cum Deep Freezer

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Refrigerator cum Deep Freezer |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Refrigerator cum deep freezer maintains two distinct temperature zones. The refrigerator zone is for chilling above zero and freezer zone is for sub-zero temperatures. |
| 2 | Operational Requirements |
| 2.1 | Fridge is required at temperatures +5 ℃ to +15 ℃ and Freezer to maintain -10 ℃ to - 20 ℃. |
| 3 | System Configuration |
| 3.1 | Refrigerator cum Deep Freezer with voltage stabilizer, complete unit. |
| 4 | Technical Specifications |
| 4.1 | Storage Capacity/Volume (approx.): Fridge: 150-260 litres; Freezer: 100-150 litres. |
| 4.2 | Construction: - Internal: Stainless steel, at least 22g. - External: Corrosion resistance (CR at least 1mm thickness). - Chest type with CFC free insulation - Solid door with lock & keys - Upright trays |
| 4.3 | Type: Compression cycled, CFC-free refrigerant preferably R-134a (both for refrigeration and insulation). |
| 4.4 | Compressor starting at 22% below rated voltage (both hot and cold starts). |
| 4.5 | Individual display for temperature inside the freezer and the fridge. |
| 4.6 | Individual alarm for Low/High temperature inside freezer and the fridge. |
| 4.7 | Target holdover time must be 15 hours or more in a continuous external temperature of 43 ℃ and 40 hours or more in a continuous external temperature of 32 ℃. |
| 4.8 | Provision for drainage for the waste water. Easy access to this waste water container for disposal of waste water. |
| 4.9 | Shall have adjustments for uneven bases. The adjustments must be easy to use like rotating a screw at the legs in the base. |
| 4.10 | Spill proof adjustable shelves/drawers. Bidder shall specify the number of shelves/drawers to be offered. |
| 4.11 | Control panel with digital display. |
| 4.12 | Humidity controller in both the compartments. |
| 4.13 | Frost free system. |
| 4.14 | Internal illumination. |
| 5 | Accessories, spares and consumables |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220 – 240 VAC, 50Hz Single Phase fitted with appropriate plug. The power cable must be at least 3 metre in length. |
| 6.3 | Shall supply suitable voltage corrector/stabilizer with the unit. |
| 7 | Standards, Safety and Training |
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 AND |
| 7.2 | CE or USFDA approved product certificate. |
| 7.3 | Shall meet IEC 60335-1 and -2-24 electrical safety standards for refrigerators and freezers. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |