Steam Autoclave

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Steam Autoclave |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Steam Sterilizers or Autoclaves are required to provide sterile systems or objects under high temperature and pressured steam. |
| 2 | Operational Requirements |
| 2.1 | Suitable for hospital dressings, linen, surgical instruments, glassware etc. |
| 3 | System Configuration |
| 3.1 | Steam autoclave, capacity approx. 90 litres, complete with accessories. |
| 4 | Technical Specifications |
| 4.1 | Single door high pressure steam sterilizer with triple wall, steam jacket and separate boiler. |
| 4.2 | Shall have advanced microprocessor control system with 128x128 pixels LCD display. |
| 4.3 | Material of construction: - Sterilizer chamber SS 316 - Door SS 316 - Jacket MS - Loading carriage SS 316 - Transfer trolley: MS, epoxy powder coated - Door Gasket: Silicon or better - Insulation: fibre glass resin bonded wool or better - Insulation cover: SS sheets |
| 4.4 | Chamber capacity: approx.90L. |
| 4.5 | Operating temperature 121 ℃- 138 ℃pressure 1.1 to 2.2 kg/ cm2 of steam pressure. |
| 4.6 | Sterilizer must be provided with steam generator. |
| 4.7 | Spring loaded safety valves and automatic vacuum breaker for jacket. |
| 4.8 | Removable plug screen for chamber drain. |
| 4.9 | SS baffle for even steam distribution in the chamber. |
| 4.10 | Safety lock for door: Pressure lock safety precautions for lid opening. |
| 5 | Accessories, spares and consumables |
| 5.1 | Accessories: - Three drums sterilising stainless steel with seamless construction suitable to fit into the steriliser. - 1 x spare heating element - A minimum of two spare lid gaskets |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220-240VAC Single phase, 50Hz, /440 V 3 Phase as appropriate fitted with appropriate plug. The power cable must be at least 3 metres long. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |