Water Bath

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Water Bath |
|  | Manufacturer |
|  | Brand |
|  | Type / Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Water bath maintains a constant pre-set temperature for treating samples. |
| 2 | Operational Requirements |
| 2.1 | General purpose water bath is required. |
| 3 | System Configuration |
| 3.1 | Water Bath, complete unit with all standard accessories. |
| 4 | Technical Specifications |
| 4.1 | Material: Inner and outer jacket made up stainless steel of approximately1mm thickness. |
| 4.2 | Capacity: approximately 20 L. |
| 4.3 | No. of coils: 2 minimum. |
| 4.4 | Microprocessor controlled programmable, digital display for temperature etc. |
| 4.5 | Temp. Range: 10 ºC to 100 ºC accuracy +/- 1 ℃. |
| 4.6 | Must have a stirrer for circulation. |
| 4.7 | Vessel Holders: Circular rings at the tops, size ramping 2 inch to 6 inch, minimum 10 nos. holding beakers of 25ml to 500ml capacity. |
| 4.8 | No. of holes: 12 holes with 75 mm concentric rings. |
| 5 | Accessories, spares and consumables |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
| 7.2 | CE (EEC Directives) or USFDA approved product certificate. |
| 7.3 | Shall meet IEC 61010-1 safety requirements for electrical equipment for measurement, control, and laboratory use. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | Supplier must accomplish proper commissioning of equipment onsite. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |