Hot Pack Unit

| S.N. | Purchaser’ Specifications |
| --- | --- |
|  | Hot Pack Unit |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Hot pack unit is used to provide comfortable and effective moist heat therapy to the patients. |
| 2 | Operational Requirements |
| 2.1 | The unit shall be mobile and must be able to run 24 hours/day. |
| 3 | System Configuration |
| 3.1 | Hot Pack Unit, complete set. |
| 4 | Technical Specifications |
| 4.1 | Fully stainless steel unit with heavy duty castor wheels with option to easily change water and heating capacity of up to 12 large packs at one time and temperature pre-set at 167 oF. |
| 4.2 | It shall have PVC coated and low-water cut off feature with special insulate to preserve heat and help conserve power consumption. |
| 4.3 | Shall come with built in adjustable thermostats and on/off switch with indicator. |
| 4.4 | It shall have concealed elements, coved bottoms, and coated racks. |
| 4.5 | It shall have mesh and hangers inside the unit to hold the pads for proper stable heating transfer. |
| 4.6 | It shall come with different sizes white clay pads and able to hold temperature up to 30 minutes of deep moist heat. |
| 4.7 | Shall come with durable custom size and shapes terry covers for hot packs with different sizes for cervical, lower back, extremities and more. |
| 5 | Accessories, spares and consumables |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 AND |
| 7.2 | CE or USFDA approved product certificate. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | Supplier must accomplish proper installation & commissioning of the equipment onsite. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English |
| 12.2 | Service (Technical / Maintenance) manual in English |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |