Interferential Therapy, Mobile with Trolley

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Mobile Interferential Therapy with Trolley |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type/Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | A current therapy used in the treatment of circulatory disorders, range of motion, oedema and muscle spasms. Interferential current is a form of electrical therapy that delivers currents to deep tissues through the use of kilohertz-carrier-frequency pulsed or sinusoidal currents to overcome the impedance offered by the skin. It is a deeper form of TENS. |  |
| 2 | Operational Requirements |  |
| 2.1 | A choice of two or four pole treatment and have a facility to enable the user to set the "beat" frequency according to the condition being treated with rechargeable internal battery. |  |
| 3 | System Configuration |  |
| 3.1 | Mobile Interferential Therapy with trolley and complete with accessories. |  |
| 4 | Technical Specifications |  |
| 4.1 | Must have low & medium frequencies current for electrotherapy. |  |
| 4.2 | 2 & 4 pole with dipole vector field with TENS. |  |
| 4.3 | Galvanic, Faradic MF surge & NME stimulation. |  |
| 4.4 | Large programmable memory with pre-set programme. |  |
| 4.5 | Carrier wave frequency adjustable between 2-10 KHz. |  |
| 4.6 | Large LCD display for treatment parameter & option of CC/CV mode. |  |
| 4.7 | Must be operable without mains power for at least 2 hours of active use. |  |
| 4.8 | Must come with rechargeable battery and adapter. |  |
| 4.9 | Must come with trolley on 4 swivel castors, of which 2 with breaks. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND |  |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
| 7.3 | Shall meet IEC-60601-1-2:2001General Requirements of Safety for Electromagnetic Compatibility. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |