Moist Heat Therapy Unit (Hydrocollator Unit)

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Moist Heat Therapy Unit (Hydrocollator Unit) |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type/Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | A Hydrocollator unit is a heating unit used to heat up the hydrocollator pack. A hydrocollator pack is a silica gel filled pack that is soaked in hot water (in a hydrocollator unit) to provide prolonged moist heat. These hydrocollator packs are used to give heat therapy to the patients. |  |
| 2 | Operational Requirements |  |
| 2.1 | The unit must be able to run 24 hours/day with a temperature from 50 0C to 60 0C or more. |  |
| 3 | System Configuration |  |
| 3.1 | Moist Heat Therapy Unit (Hydrocollator Unit), complete unit. |  |
| 4 | Technical Specifications |  |
| 4.1 | Must be made of heavy gauze stainless steel sheet, double walled and fully insulated in between. |  |
| 4.2 | Must be fitted with a 2000 watts immersion heater, pilot light, and thermostat for heat control. |  |
| 4.3 | Shall come with: - 8 Large Steam Packs: 5 standard size, 1 half, 1 neck and 1 full back size. - 4 Small Packs: standard size |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3m long. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 AND |  |
| 7.2 | CE or USFDA approved product certificate. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 1 year. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | Supplier must accomplish proper installation & commissioning of the equipment onsite. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English |  |
| 12.2 | Service (Technical / Maintenance) manual in English |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |