Muscle Stimulator, Electrical

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Muscle Stimulator, Electrical |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type/Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | Electrical muscle stimulation is defined as the use of an electrical current to transfer energy to a wound or to a surface for relieving pain. |  |
| 2 | Operational Requirements |  |
| 2.1 | Electrical muscle stimulator to generate electrical pulses to a surface for relieving pain and to restore or improve their function. |  |
| 3 | System Configuration |  |
| 3.1 | Electrical Muscle Stimulator complete with electrodes and other accessories. |  |
| 4 | Technical Specifications |  |
| 4.1 | Digital output meter. |  |
| 4.2 | Galvanic, intermittent galvanic, faradic, surge faradic. |  |
| 4.3 | Pulse duration 0.01, 0.03, 0.1, 0.3, 1, 3, 10, 30, 100, 300m second. |  |
| 4.4 | Pulse rest time 1-100 pulse/second variable. |  |
| 4.5 | Visual & audio indicator. |  |
| 4.6 | Faradic frequency 50 Hz/0.7 msec. |  |
| 4.7 | Surge duration 1-6 sec. continuously adjustable. |  |
| 4.8 | Digital Timer. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Two rubber electrodes with leads. - One Pin electrode with lead. |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
| 7.3 | Shall meet IEC 60601-2-10 Safety requirements - Part 2: Particular requirements for the safety of Nerve and Muscle Stimulators. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 1 year. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation, Inspections and Commissioning |  |
| 11.1 | Supplier must accomplish proper installation & commissioning of the equipment on site. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English |  |
| 12.2 | Service (Technical / Maintenance) manual in English |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |