Traction Bed

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Traction Bed |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type/Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | Traction bed is used for applications of traction therapy. |  |
| 2 | Operational Requirements |  |
| 2.1 | Manually operated traction bed. |  |
| 3 | System Configuration |  |
| 3.1 | Traction bed, complete unit. |  |
| 4 | Technical Specifications |  |
| 4.1 | It shall have facility to fit any traction unit. |  |
| 4.2 | It shall be a four section traction bed. |  |
| 4.3 | Dimension (approx.): 190x75x75cm |  |
| 4.4 | The lumbar section shall roll freely on rollers for friction-free lumbar pull. |  |
| 4.5 | Cervical traction angle is controlled by vertical adjustment of the machine mounting board. |  |
| 4.6 | It shall have facility to adjust the height of the traction unit. |  |
| 4.7 | It shall have armpit for counter traction |  |
| 4.8 | Head section shall be adjustable from 0° to 90° |  |
| 4.9 | It shall be fitted with retractable wheels. |  |
| 4.10 | The bed shall be epoxy powder coated washable paint finish. |  |
| 4.11 | It shall come with mattress 5cm. |  |
| 4.12 | The lifting capacity of traction bed shall be 180kg or more. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | Accessories: - Flexion stool – 01 no. - Lumbar & Cervical harnesses – 01 no. - Spreader bar – 01 no. - Cervical Pillow and Cover – 01 no. |  |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Climate, Temperature, Humidity, etc. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND |  |
| 7.2 | CE or USFDA approved product certificate. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 1 year after acceptance. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |