Transcutaneous Electrical Nerve Stimulator (TENS)

| S.N. | Purchaser’s Specifications |
| --- | --- |
|  | Transcutaneous Electrical Nerve Stimulator |
|  | Manufacturer |
|  | Brand |
|  | Type/Model |
|  | Country of Origin |
| 1 | Description of Function |
| 1.1 | Transcutaneous Electric Nerve Stimulator (TENS) is a method of treating pain that is non-invasive and does not use pharmaceuticals. |
| 1.2 | The TENS device sends impulses through the skin that stimulate the nerve (or nerves) in the treatment area. In many cases this stimulation will greatly reduce or eliminate the pain sensation felt by masking the original pain message sent to the brain. |
| 2 | Operational Requirements |
| 2.1 | Transcutaneous Electrical Nerve Stimulation (TENS) is indicated to be used under a physician’s prescription for the symptomatic relief and management of chronic (long term) pain and for the treatment of postoperative or posttraumatic pain. |
| 3 | System Configuration |
| 3.1 | The system consists of: - Transcutaneous electric nerve stimulator - LCD display - Timer - Rechargeable battery |
| 4 | Technical Specifications |
| 4.1 | Microprocessor /Microcontroller circuit. |
| 4.2 | Programmable LCD display with easy read-out. |
| 4.3 | Simultaneous separate treatment to two patients. |
| 4.4 | More than 350 programmable output currents. |
| 4.5 | Feather touch key operations. |
| 4.6 | Programmable 0 – 99 min count down timer with LCD display. |
| 4.7 | Inbuilt system ensures auto switch OFF of the unit when treatment time elapses. |
| 4.8 | Separate output for Galvanic (G/IG/TNS) and Faradic (F/SF) currents. |
| 4.9 | Galvanic Current DC. |
| 4.10 | 31 user selectable pulse durations for IG. |
| 4.11 | 11 user selectable pulse frequencies for IG. |
| 4.12 | TENS output selectable from three modes. |
| 4.13 | 3 modes of TENS (Transcutaneous Electrical Nerve Stimulator) includes: - Continuous Mode: Continuous pulse trains are delivered with user selectable pulse rate from 20 to 125 pulses per second. - Burst Mode: 2 bursts of 10 pulses are emitted per second with pulse width of 200usec. - Modulated Mode : Pulse trains are delivered for 2 Sec and rest for 1 sec. |
| 4.14 | Faradic / Surge faradic output facility. |
| 4.15 | Surge duration selectable in steps of 0.5 sec. |
| 4.16 | Surge delivered : selectable 0 to 9.5 sec |
| 4.17 | Inbuilt rechargeable battery for at least one hour operation. |
| 5 | Accessories, spares and consumables |
| 5.1 | Accessories: - Colour coded leads-4 no. - Rectangular Electrodes-4 no. - Sponges mains lead-4 no. - Rechargeable battery-1 no. - Jelly 1 bottle ( 250 ml ) |
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |
| 6 | Operating Environment |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |
| 6.3 | SMPS power management technology and spike suppressor. |
| 7 | Standards and Safety Requirements |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |
| 7.3 | Shall meet IEC 60601-2-10 Safety requirements - Part 2: Particular requirements for the safety of Nerve and Muscle Stimulators. |
| 8 | User Training |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |
| 9 | Warranty |
| 9.1 | Comprehensive warranty for 2 years. |
| 10 | Maintenance Service During Warranty Period |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |
| 11 | Installation and Commissioning |
| 11.1 | Supplier must accomplish proper installation & commissioning of the equipment on site. |
| 12 | Documentation |
| 12.1 | User (Operating) manual in English. |
| 12.2 | Service (Technical / Maintenance) manual in English. |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |
| 12.4 | Certificate of calibration and inspection from factory. |