Ultrasound Therapy Unit (Single Head)

| S.N. | Purchaser’s Specifications |  |
| --- | --- | --- |
|  | Ultrasound Therapy Unit (Single Head) |  |
|  | Manufacturer |  |
|  | Brand |  |
|  | Type/Model |  |
|  | Country of Origin |  |
| 1 | Description of Function |  |
| 1.1 | Ultrasound uses a high frequency sound wave emitted from the sound head when electricity is passed through a quartz crystal. The sound waves cause the vibration of water molecules deep within tissue causing a heating effect. When the sound waves are pulsed, they cause a vibration of the tissue rather than heating. The stream of sound waves helps with nutrition exchange at the cellular level and healing. Ultrasound is helpful for ligament healing and clinically, for carpal tunnel syndrome, and muscle spasm. |  |
| 2 | Operational Requirements |  |
| 2.1 | Single head Ultrasound therapy unit, easy to use. |  |
| 3 | System Configuration |  |
| 3.1 | Ultrasound Therapy Unit (Single Head) complete with accessories. |  |
| 4 | Technical Specifications |  |
| 4.1 | Microprocessor based with Continuous & Pulsed modes. |  |
| 4.2 | Output power 15W in continuous mode and 21W in pulse mode. |  |
| 4.3 | Pulse frequency 100Hz. |  |
| 4.4 | Output frequency 1 MHz. |  |
| 4.5 | Timer 0-15 minutes, pre-settable. |  |
| 4.6 | Time adjustment up to 99 minutes. |  |
| 4.7 | Two digital display meters indicate the output in w/cm². |  |
| 4.8 | Patient safety circuit. |  |
| 5 | Accessories, spares and consumables |  |
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| 6 | Operating Environment |  |
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. |  |
| 7 | Standards and Safety Requirements |  |
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND |  |
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate. |  |
| 7.3 | Shall meet IEC-60601-2-5 Part 2-5: Particular requirements for the safety of Ultrasonic Physiotherapy equipment. |  |
| 8 | User Training |  |
| 8.1 | Must provide user training (including how to use and maintain the equipment). |  |
| 9 | Warranty |  |
| 9.1 | Comprehensive warranty for 2 years. |  |
| 10 | Maintenance Service During Warranty Period |  |
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required. |  |
| 11 | Installation and Commissioning |  |
| 11.1 | Supplier must accomplish proper installation & commissioning of the equipment onsite. |  |
| 12 | Documentation |  |
| 12.1 | User (Operating) manual in English. |  |
| 12.2 | Service (Technical / Maintenance) manual in English. |  |
| 12.3 | List of important spare parts and accessories with their part numbers and costing. |  |
| 12.4 | Certificate of calibration and inspection from factory. |  |