Technical Specification of De-ionised Water Plant

	S.N.
	Purchaser’s Specifications
	Bidder’s compliance sheet

	
	De-ionised Water Plant
	Compliance
(Yes/No)
	Corresponding page no. of submitted data sheet/ catalogue in support of specification
	Remarks

	
	Manufacturer
	
	
	

	
	Brand
	
	
	

	
	Type/Model
	
	
	

	
	Country of Origin
	
	
	

	
	Name of Bidder
	
	
	

	1
	Description of Function
	
	
	

	1.1
	The Deionized Water / Water Purification Plant is intended to produce high-purity water by removing dissolved ions and impurities from potable water, suitable for use in laboratory, clinical, and analytical applications.
	
	
	

	2
	Operational Requirements
	
	
	

	2.1
	De-ionised water plant with a capacity of at least 10 litres/hour or higher.
	
	
	

	3
	System Configuration
	
	
	

	3.1
	De-ionised Water Plant supplied as a complete system with all standard accessories required for installation, operation, and commissioning.
	
	
	

	3. 2
	The system shall include a storage tank with a CO2-absorbing vent filter.
	
	
	

	4
	Technical Specifications
	
	
	

	4.1
	Flow rate: Minimum 10 litres per hour or higher
	
	
	

	4.2
	Resistivity: 10-15 MΩ·cm or better. 
	
	
	

	4.3
	TOC: < 30 μg/lt 
	
	
	

	4.4
	Bacteria: <1cfu/ml.
	
	
	

	4.5
	Water Conductivity:  1 to 2 μS/cm or less 
	
	
	

	4.6
	The system must be supplied with a Reverse Osmosis (RO) plant of appropriate capacity to serve as the primary purification stage for the input of the DI (De-ionisation) stage.
	
	
	

	4. 7
	Must be supplied with continuous resistivity and/or conductivity meters to monitor the performance of the De-ionised Water Plant.
	
	
	

	5
	Accessories, Spares and Consumables
	
	
	

	5.1
	All standard accessories, spares, and consumables required for installation, operation, and routine maintenance of the De-ionised Water Plant shall be included in the offer. Bidders shall clearly specify the quantity of each item supplied, including items not explicitly mentioned in the technical specifications.
	
	
	

	6
	Operating Environment	
	
	
	

	6.1
	The system offered shall be designed to store and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
	
	
	

	6.2
	Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs to meet purchaser’s country requirements. The power cable must be minimum 3 metres long. 
	
	
	

	7
	Standards and Safety Requirements
	
	
	

	7.1
	Must submit valid ISO 9001 or ISO 13485:2016 (or latest) certification.
	
	
	

	7.2
	Must submit valid CE or USFDA, approved certification.
	
	
	

	8
	User Training
	
	
	

	8.1
	The supplier shall provide on-site user training covering operation, routine maintenance, and basic troubleshooting of the De-ionised Water Plant at the time of installation and commissioning.
	
	
	

	9
	Warranty
	
	
	

	9.1
	The bidder must provide a full comprehensive warranty (covering parts and labor) for a period of 02 years starting from the date of formal acceptance/commissioning
	
	
	

	10
	Maintenance Service During Warranty Period
	
	
	

	10.1
	During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
	
	
	

	10.2
	In the event of a breakdown, the bidder must provide technical support or on-site service within 24 hours of notification.
	
		
	

	11
	Installation and Commissioning 
	
	
	

	11.1
	The supplier shall ensure that the De-ionised Water Plant is installed and commissioned by certified or qualified personnel. Any prerequisites for installation shall be communicated to the purchaser in advance, in detail, and the unit shall be fully operational upon completion.
	
	
	

	12
	Documentation
	
	
	

	12.1
	Bidder must submit a valid authorization from the manufacturer.
	
	
	

	12.2
	Certificate of Calibration and Inspection issued by the manufacturer’s factory. This certificate must correspond to the specific serial number of the equipment delivered.
	
		
	

	12.3
	User (Operating) manual in English.
	
	
	

	12.4
	Service (Technical / Maintenance) manual in English. 
	
	
	

	12.5
	List of important spare parts and accessories with their part number and costing.
	
	
	

	Note: Bidders must completely fill the Technical Specification Form (TSF). Only 'Yes/No' remarks are insufficient. The specific page number in the submitted data sheet/ catalogue for every required parameter must be clearly mentioned and highlighted. Failure to provide this evidence may lead to the rejection of the bid from the technical committee.



